Instructions for Use (IFU)

Labeling Documentation — PulseTrack Case Study

NOTE: PulseTrack is a fictional device used for regulatory affairs training purposes. This document is a portfolio exercise, not a real regulatory filing.

# 1. Purpose

The Instructions for Use (IFU) is both a usability document and a documented risk control under ISO 14971 — warnings and precautions listed here are formal 'information for safety' risk controls, traceable back to the Risk Management File, not just marketing or convenience text.

# 2. Device Description

PulseTrack is a wireless fingertip pulse oximeter that measures blood oxygen saturation (SpO2) and pulse rate, transmitting readings via Bluetooth Low Energy to the PulseTrack companion mobile application.

# 3. Indications for Use

For monitoring SpO2 and pulse rate in adult patients (18 years and older) with a physician-diagnosed respiratory condition placing them at risk of hypoxemia, in home and ambulatory non-clinical settings, as an adjunct to — not a replacement for — clinical judgment and physician-directed care.

# 4. Warnings

• Do not use as the sole basis for diagnosis or treatment decisions.
• Do not use during transport, surgery, or in critical care environments.
• Seek immediate medical attention if you experience symptoms of low oxygen (shortness of breath, confusion, bluish lips or fingertips) regardless of the device reading.
• Nail polish, artificial nails, or poor peripheral circulation may cause inaccurate readings.

# 5. Precautions

• Ensure the sensor is fully seated on a clean, dry finger.
• Keep your phone within 10 meters (33 feet) of the device to maintain Bluetooth connection.
• Charge the device fully before first use; a low-battery alert will appear at 20% and 10% remaining charge.
• Rotate finger placement or remove the device periodically during extended monitoring sessions to avoid skin irritation.

# 6. Instructions for Use — Step by Step

1. Charge PulseTrack fully via the included USB-C cable.
2. Download and open the PulseTrack companion app; pair via Bluetooth following on-screen prompts.
3. Clip the sensor onto a clean, dry fingertip, ensuring the indicator light faces upward.
4. Wait 10 seconds for the reading to stabilize; the app will display SpO2 and pulse rate.
5. If the app displays a 'low signal quality' indicator instead of a numeric reading, reposition the sensor or try a different finger.
6. Configure your low-SpO2 alert threshold in the app settings as directed by your healthcare provider.

# 7. Maintenance and Cleaning

Wipe the sensor with a soft cloth dampened with isopropyl alcohol (70%) between uses. Do not immerse in water. Do not sterilize by autoclave.

# 8. Electronic IFU (eIFU) Note

This IFU content is also maintained in a structured XML format (see PulseTrack_IFU_Structured_XML_Sample.xml) to support electronic/multi-language labeling distribution consistent with current EU eIFU practices, enabling machine-readable content management across markets and translations without re-authoring the underlying text for each language version.